| Service Manual | Data Communication | Doc. No |  |
| --- | --- | --- | --- |
| 1. Information for data communication Many of TOMEY products can send measurement data to a PC. This is an instruction to do data communication between PC and TOMEY units through a serial port. Required tools - RS-232C cable (crossover type) - Gender changer (depends on type of RS-232C cable) - Software (ex: Hyper terminal) Gender changer Communication protocol X-modem: TL-900, TL-2000, TL-3000, TL-2000A, TL-3000A, TL-3000B, TR-1000, RC-1000, TR-3000, RC-3000, TR-4000, RC-4000, AL-2000, UD-6000 (older software than 1.30/2.20), RT-6000 (Ref data), RC-5000. Z-modem: AL-100, SP-3000, AL-3000, UD-6000 (new software than and 1.30/2.20), SP-100. Other way: TMS-3, TMS-4, RT-6000(Topo data) <Page 1> |  |  |  |

| Service Manual | Data Communication | Doc. No |  |
| --- | --- | --- | --- |
| 2. Communication with X-modem format product. - Connect the Computer and TOMEY unit by a RS-232C cable. (If you need, use a gender changer) - Open software which can read binary code. (For example: HyperTerminal) - Adjust communication condition on Tomey unit and communication software of the computer (Baud rate, Data bit, Parity, Stop bit and so on) - Collect some data on the Tomey unit and send data to the computer (Carriage return should be active by setting of software) *Caution If Tomey product cannot change communication condition, PC has to adjust as follows. Baud Rate; 2400 bps. Data bit; 7bit. Parity; Even. Stop bit; 1bit. Data code; ASCII. Flow control; Xon/Xoff. <Page 2> |

| Service Manual | Data Communication | Doc. No |  |
| --- | --- | --- | --- |
| 3. Communication with Z-modem format product. - Connect the Computer and TOMEY unit by a RS-232C cable. (If you need, use gender changer) - Open software which can receive Z-mode-data. (For example: HyperTerminal) - Adjust communication condition on Tomey unit and communication software of the computer. (Baud rate, Data bit, Parity, Stop bit and so on) - Collect some values on the Tomey unit and send data to the computer ### Information - You will send measurement data with “CSV” format. This file can be read with Microsoft Excel. - File name is saved as “TOMEY. csv”. - If you select simple Z-modem for receiving data it will save the last file as TOMEY.CSV, the computer automatically renames the older data to TOMEY0.CSV TOMEY1.CSV TOMEY2.CSV ... - If you select Z-modem with crash recover, new data is also “TOMEY. CSV” file. It must be removed for the next measurement <Page 3> |  |  |  |